Sterile Hold Validation
Helper (test version)

User guide

# Interface
Fig.1. Program User Interface
1. “Load file” button – opens window to let user choose file. At this moment program accepts *.xlsx Excel files. The loaded data are presented at “Data” tab (see Fig.2). Table is interactive, so user can navigate through data. After loading data, a number of termocouples (TC’s) is displayed and overall process time (timestamp from last low from file subtract timestamp from first row) is displayed.
2. “Validate” button – program analyse data loaded from file and finds the time at which last TC reaches 121.0ºC, which starts Sterile Hold Regime (SHR) and displays it as T1, finds the time at which first TC reaches 121.0ºC on SHR, which leads to end of Regime and displays it as T2. ΔT=|T1-T2| is calculated and displayed on the screen. If ΔT is in range between 15 minutes and 15 minutes 30 seconds, the ΔT is coloured as green. If ΔT is above or below that range, it will be coloured as red. The data in the table are “cut” and only data from SHR are displayed (see Fig.2). If during SHR any of TC does not drop below 121.0ºC or rise above 123.0ºC, then the message “Sterile Hold valid!” indicator will have green background. If this rule is not met, then “Sterile Hold invalid!” indicator will have green background.
3. “Post-calculate” button – this button triggers number the plotting sequences and number of calculations. In the “F_zero” Post-Calculation tab the plot of SHR is displayed. The average Value of F0 is calculated for whole SHR.
Fig.2. “Data” tab and data representation from file
Fig.3. Effect of “Validate” button